| Risk Management Plan | Supplemental Risk Management Plan for the component subassembly, revision 1 |
| --- | --- |

| Scope | This document will establish a supplemental Risk Management Plan for the component subassembly (manufactured for supplier) as defined within the Quality System Agreement supplier-manufacturer. It is a supplement to the overall Risk Management Plan for the medical device system. |
| --- | --- |
| Risk Management Process, Acceptance Criteria | This Risk Management Plan is generated in accordance to ISO 14971:2019 and in alignment to SOP Risk Management (supplier) and SOP Risk Management (manufacturer). As different legal entities and therefore different Quality Management Systems will revert to this Risk Management Plan, definitions will be taken from ISO 14971:2019. Definitions and data sources applied, are listed in Annex 1 to this Risk Management Plan. Ratings used for severity of harm are given in annex 2. Ratings used for occurrence are given in annex 3. Ratings used for detectability are given in annex 4. Risk acceptability criteria are given in annex 5. |
| Assignment of roles, responsibilities and authorities | To accomplish all risk management activities, a risk management team was established. The following persons from supplier were assigned as risk management team members: Production Engineer, supplier Quality Engineer, supplier Development, supplier Additional resources may be consulted as needed. The development function of the manufacturer will coordinate all activities regarding specifications and failure modes, hazardous situations and harms, severity and probability of occurrence of harm. The risk management team members are mandatory reviewer and approvers of the risk management activities within this risk management plan. |

| Risk Management Activities | Responsible person | Time plan |
| --- | --- | --- |
| Risk ANALYSIS ON System level | The Risk Analysis for the final finished medical device system is out-of-scope of this Risk Management plan and is in responsibility of manufacturer. |
| Risk ANALYSIS ON process level | A Process Risk Analysis (PFMEA) will be established assessing the risks of production of the component subassembly at supplier. | September 2021 |
| Verification of Risk Management Activities | Traceability to all production control measures will be given within the PFMEA. Review of PFMEA will include verification, that production control measures were implemented. | September 2021 |
| Risk management report | Risk Management report for the final finished medical device system is out-of-scope of this Risk Management Plan and is in responsibility of manufacturer. |
| Verification of Effectiveness of Risk Management Process | The effectiveness of the production risk control measures will be validated by continuous monitoring of production and post-production information. An initial review will be established within the approval of PFMEA. | September 2021 |
|  | The effectiveness of the Risk Management process for the final finished medical device system is out-of-scope of this Risk Management Plan and is in responsibility of manufacturer. |
| collection and review of relevant production and post-production information | Supplier will collect and evaluate production and post-production information as defined in supplier Quality Manual | n/a, continuous activities |
|  | The collection and evaluation of production and post-production information for the final finished medical device system is out-of-scope of this Risk Management Plan and is in responsibility of manufacturer. |
| RISK MANAGEMENT FILE | The Risk Management file for the final finished medical device system is out-of-scope of this Risk Management plan and is in responsibility of manufacturer. |

| Approval |
| --- |
| Function | name / date / signature |
| Production Engineer, Supplier |  |
| Quality Engineer, Supplier |  |
| Development, Manufacturer |  |
| Quality Management Representative, Supplier |  |
| Quality Management Representative, Manufacturer |  |

| Annex 1 | Definitions and source of data |
| --- | --- |

Definitions, where possible, will be taken from ISO 14971:2019

| Term | Definition | Source of data | Notes |
| --- | --- | --- | --- |
| Harm | injury or damage to the health of people, or damage to property or the environment | Manufacturer as owner of the final finished product will determine the applicable list of harms | Approval of this risk analyses will document manufacturer approval of list of harms |
| Hazard | potential source of harm | Baseline for a list of hazards is given within table C.1 of ISO 14971:2019 | - |
| Hazardous situation | circumstance in which people, property, or the environment are exposed to one or more hazard(s) | Manufacturer as owner of the final finished product will determine the applicable list of hazardous situations | Approval of this risk analyses will document Manufacturer approval of list of hazardous situations |
| Risk | combination of the probability of occurrence of harm and the severity of that harm | - | - |
| Severity | measure of the possible consequences of a hazard | Manufacturer as owner of the final finished product will determine the applicable severity ratings | Approval of this risk analyses will document Manufacturer approval of list of severity ratings. Severity ratings are listed in Annex 2 of this Risk Management Plan. |
| Occurrence of harm | probability of occurrence of harm | Manufacturer as owner of the final finished product will determine the applicable occurrences | Approval of this risk analyses will document Manufacturer approval of occurrences of harm. Occurrence ratings are listed in Annex 3 of this Risk Management Plan |
| Occurrence of failure | probability of occurrence of failure | Supplier will establish the occurrence of failure based on historical process data | Occurrence of failure is to be differentiated from occurrence of ham. Occurrence ratings are listed in Annex 3 of this Risk Management Plan. |
| Detectability of failure | probability of detection of failure before the final finished product will be placed on the market | Supplier will determine the detectability of failure based on historical process data | Detectability ratings are listed in Annex 4 of this Risk Management Plan. |

| Annex 2 | Severity ratings |
| --- | --- |

As different legal entities and therefore different Quality Management Systems will revert to this Risk Management plan, an alignment of two different systems for severity ratings is needed.

The following table shows the alignment of both systems.

| Supplier | Manufacturer |
| --- | --- |
| Severity | Description | Severity | Description |
| 10 |  | 5 |  |
| 9 |  | 4 |  |
| 8 |  |  |  |
| 7 |  | 3 |  |
| 6 |  | 2 |  |
| 5 |  | 1 |  |
| 4 |  |  |  |
| 3 |  |  |  |
| 2 |  |  |  |
| 1 |  |  |  |

| Annex 3 | Ratings for Occurrence |
| --- | --- |

As different legal entities and therefore different Quality Management Systems will revert to this Risk Management Plan, an alignment of occurrence ratings is needed.

- Manufacturer will apply occurrence ratings for harm

  - These occurrence ratings are to be understood as probability of occurrence of harm.

- Supplier will apply occurrence ratings for failure

  - These occurrence ratings are to be understood as occurrence of failure.

  - Probability of occurrence of failure is to be communicated as such to manufacturer.

The following table shows the alignment of both systems.

| Supplier | Manufacturer |
| --- | --- |
| Occurrence | Description | Occurrence | Description |
| 10 |  | 5 |  |
| 9 |  |  |  |
| 8 |  |  |  |
| 7 |  | 4 |  |
| 6 |  |  |  |
| 5 |  |  |  |
| 4 |  | 3 |  |
| 3 |  |  |  |
| 2 |  |  |  |
| 1 |  | 2 |  |
|  |  | 1 |  |

| Annex 4 | Ratings for Detectability |
| --- | --- |

As different legal entities and therefore different Quality Management Systems will revert to this Risk Management plan, an alignment of detectability ratings is needed.

- Manufacturer will not apply detectability ratings.

- Supplier will apply detectability ratings

  - These detectability ratings are to be understood as detectability of the failure.

The following table shows the alignment of both systems.

| Supplier | Manufacturer |
| --- | --- |
| Detect-ability | Description | Detect-ability | Description |
| 10 |  | 5 | n/a |
| 9 |  | 4 | n/a |
| 8 |  | 3 | n/a |
| 7 |  | 2 | n/a |
| 6 |  |  |  |
| 5 |  |  |  |
| 4 |  |  |  |
| 3 |  |  |  |
| 2 |  | 1 | n/a |
| 1 |  |  |  |

| Annex 5 | Risk Acceptance Criteria |
| --- | --- |

As manufacturer is responsible for the final finished medical device system, the applicable risk acceptance criteria from manufacturer will be applied.